## Slide 1: Spectrometry_Gamma A New Module For The EduGATE Project
Uwe Pietrzyk & Michaela Gaens
Institute of Neurosciences and Medicine (INM), Research Center Juelich, Germany
Department of Mathematics and Natural Sciences, University of Wuppertal, Germany
Contact: u.pietrzyk@fz-juelich.de or uwe.pietrzyk@uni-wuppertal.de
(Basic Paper  „Zeitschrift für Medizinische Physik“ / Z. Med. Phys. 23 (2013) 65-70)
Notes: <number>

## Slide 2
Long is the way of theory,
short and effective by examples –
(Lucius Annaeus Seneca (the Younger),
Epistulae morales)
Notes: <number>

## Slide 3
  Spectrometry_Gamma:
    introduces to the basics of gamma spectra obtained from a scintillation detector
    Inspired by chapter 10 of Physics in Nuclear Medicine (Cherry, Sorensen & Phelps, 2012)
    provides GATE-macros
    comes in two versions: Basic & Extended
    Programme Code in C and for ROOT to calculate and plot relevant parameters during Compton Scatter
Notes: <number>

## Slide 4
Pulse-height spectrum for 137Cs showing
typical structures
from: Cherry et al, 2012
Notes: <number>

## Slide 5
Basics of
Compton_Scatter
Esc = E0 /[1 + ( E0 /0.511) (1 − cos)]
Ere = E0 − Esc
for 180-degree scattering:
Escmin = E0 /[1 + (2 E0 /0.511)]
Eremax = E0 – Escmin
(Esc)
(E0)
(Ere)
    (Cherry, et al, 2012)
Notes: <number>

## Slide 6
Code
for (int i=0; i<n_values; i++)
{
// energy of the incident gamma
E_gamma[i] = (i+1) * e_step;
// energy of the recoil electron; -> "compton electron" -> "compton edge"
Emax_re[i] = E_gamma[i] * E_gamma[i] / (E_gamma[i] + m_electron_2);
// energy of the back-scattered gamma
Emin_sc[i] = E_gamma[i] - Emax_re[i];
printf( " E_gamma: %f Emax_re: %f Emin_sc: %f (MeV) \n", E_gamma[i], Emax_re[i], Emin_sc[i]);
}
--> output
E_gamma: 0.010000 Emax_re: 0.000377 Emin_sc: 0.009623 (MeV)
E_gamma: 0.020000 Emax_re: 0.001452 Emin_sc: 0.018548 (MeV)
E_gamma: 0.030000 Emax_re: 0.003152 Emin_sc: 0.026848 (MeV)
E_gamma: 0.040000 Emax_re: 0.005415 Emin_sc: 0.034585 (MeV)
.
.
E_gamma: 0.660000 Emax_re: 0.475806 Emin_sc: 0.184194 (MeV)
E_gamma: 0.670000 Emax_re: 0.485035 Emin_sc: 0.184965 (MeV)
E_gamma: 0.680000 Emax_re: 0.494281 Emin_sc: 0.185719 (MeV)
E_gamma: 0.690000 Emax_re: 0.503543 Emin_sc: 0.186457 (MeV)
E_gamma: 0.700000 Emax_re: 0.512820 Emin_sc: 0.187180 (MeV)
running: root -l compton_scatter_root.c
Notes: <number>

## Slide 7
Plot created via: root -l compton_scatter_root.c
Notes: <number>

## Slide 8
situation for
E = 660KeV
like for 137Cs
Notes: <number>

## Slide 9
Pulse-height spectrum for 137Cs showing
typical structures
full
energy
peak
recoil
electron
gamma,
scattered
at ~180 ° from
surrounding
material
from: Cherry et al, 2012
Notes: <number>

## Slide 10
Spectrometry_Gamma_Extended
front phantom
crystal
back phantom
Notes: <number>

## Slide 11
Extended Setup
simplified Gamma Camera  2 crystals, no collimator,
collimated gamma source
head 1
head 0
  Gamma Camera:
  Phantom in front to enforce scatter and to simulate effects of collimator material like characteristic X-rays
  Phantom in the back to enforce back-scatter
Gamma Camera  1 crystal, no collimator,
collimated gamma source (rays directed towards crystal)
two phantoms (front & back)
front phantom
gamma ray direction
back phantom
crystal
Notes: <number>

## Slide 12
situation for
E = 511KeV
E_gamma: 0.510000 Emax_re: 0.339778 Emin_sc: 0.170222 (MeV)
Notes: <number>

## Slide 13
Starting Extended Setup (Gate_v7.0)
specify:
energy &
energy resolution
in a shell-window type:
config_starter_mac_70.sh
Notes: <number>

## Slide 14
Starting Extended Setup (Gate_v7.0)
can change selection presets in Spectrometry_Gamma.txt
ViewPointThetaPhi: 0 90; 90 0; -90 0; 89 90; 15 30; 30 30; 45 45; 60 60;
VisuOnOff: novisu; visu;
SourceActivity: 100 Bq; 1000. Bq; 10000. Bq; 100000. Bq;
PhanThicknessFront: 2. mm; 0.1 mm; 0.2 mm; 0.5 mm; 1. mm; 5. mm; 10. mm; 20. mm; 30. mm;
PhantomMaterialFront: Vacuum; Water; Plexiglass; Air; Lead; Copper; Tungsten; Iodine;
PhanThicknessBack: 5. mm; 0.1 mm; 0.2 mm; 0.5 mm; 1. mm; 2. mm; 10. mm; 20. mm; 30. mm;
PhantomMaterialBack: Lead; Vacuum; Water; Plexiglass; Air; Copper; Tungsten; Iodine;
CrystalMaterial: NaI; BGO; LSO; GSO; PWO; LuAP; YAP; CZT; Silicon; Lead; Tungsten;
SourceVolMaterial: Vacuum; Water; Plexiglass; Air; PVC;
SourceType: gamma_collim; gamma;
SourceEnergy: 80; 10; 50; 60; 90; 100; 120; 140; 160; 200; 240; 364; 511; 662; 1000; 1600;
E_res: 0.10; 0.00; 0.01; 0.05; 0.15; 0.20; 0.25;
Notes: <number>

## Slide 15
note: different scales!!
full
energy
peak
(662 keV)
recoil electron 
compton edge
back-scatter peak
Notes: <number>

## Slide 16
full energy peak
(1600 keV)
Single
escape peak
Double
escape peak
The high energy of the incident photon (1600 keV) allows the production
of e+-e- -pairs. Subsequently, the positron annihilates with an electron to
produce two annihilation photons of 511 keV each. If one or both
annihilation photons escape, this gives rise for the escape peaks.
Fig 10.5 from Cherry et al, 2012
Notes: <number>

## Slide 17
full energy peak
(1600 keV)
Single
escape peak
Double
escape peak
Notes: <number>